CopiBook Calibration

2011-09-14 WED

This document assumes familiarity with CopiBook use.

Materials

- cleaning solution

- dust brush

- large white sheet, as big as the book cradle

- smaller black sheet

- the Kodak grey scale target

- sample page with lots of text and some images (this document will do fine)

- level with a 45-degree bubble

Prep

- Turn on the lights to let them warm up

- In the meantime, restart the machine,

- Use the level to adjust the angle of the lights to 45 degrees, and position them at equal distances from the machine, and

- Clean the top and bottom surfaces of the glass and dust off the book cradle below.

Focus Adjustment

- Once the machine has started up, click on  to access to the “scanner configuration” menu.

- In the “Lighting” tab, select “Deactivated”:

- Move over to the "Balance/Focus" tab.

- Place the sample page on the book cradle, and check that it is raised to the glass plate level.

- Click on the “Autofocus” button , wait for that to finish, then click the green check mark.

Black and White Balance

- Place the white sheet on the book cradle and then the black sheet partially over it, like so:

- Click the CopiBook configuration button, then click the White Balance button  and wait until the end of the calibration process. Then do this one step a second time. Then click the green check mark then the red x.

Lighting Correction

- Go back to the project selection window.

- Click on the Administration button .

- Select the Lighting tab.

- Place just the large white sheet on the book cradle (not the black sheet too), and raise it to the glass level.

- Type in the word “lighting” as the label for the lighting correction.

- Click on the  button and take care to cast no shadow on the white sheet during calibration. Wait for that to finish.

- The lighting correction must now be applied in the current project. Click the green check mark a couple of times until you can select a project, and then click on this icon:

- In the Lighting tab, select your lighting correction label: “lighting”.

Black and White Balance (Redux)

- Once the light correction is applied, go back and do the Black and White balance again, just once.

Testing

- Shoot the Kodak gray scale target on top of the large white sheet.

- Inform the imaging supervisor to move the tiff file for that target to the calibration verification folder and to inspect that image with the Photoshop ‘info’ tool (F8).